## Slide 1: Grant Saggars @ University of Kansas grant-saggars@ku.edu November 4, 2025

## Slide 2: Background & Outline
1. Basics of Conductive Atomic Force Microscopy (C-AFM)
2. Basics of carbon nanotubes (CNTs)
3. Combining the two

## Slide 3: When is C-AFM a useful technique?
Conducting Atomic Force Microscopy
1. Capacitance (SCM)
2. Tunneling Behavior (TUNA)
3. Spreading Resistance (SSRM)

## Slide 4: C-AFM Basics
Circa 1996: Record the current flowing at the tip/sample nanojunction simultaneously to the topography.
Figure 1.3 Block diagram of a conventional conductive atomic force microscope. Compared to the AFM, the three new elements are the conductive tip, preamplifier, and sample bias.
C-AFM measures current density; effective area matters.
  This is not the same as the sample contact area
The propagation of the electric field is dependent on both the tip and the material
Conductive Atomic Force Microscopy | Wiley Online Books

## Slide 5: Conduction Processes in Tip/Semiconductor Nanojunctions
Thermal excitation releases “stuck” charges
Frenkel-Poole Emission
Direct excitation of carriers over barrier
Applies to metal-insulator / insulator-semiconductor interfaces
Shottkey emission is a subclass of Thermionic
Thermionic Emission
Highly insulating, semiconducting
Occurs when tunneling through trapezoidal barrier
Generally happens in thicker samples
Fowler-Nordheim
Highly insulating, semiconducting
Occurs when tunneling through square barrier
Direct Tunneling

## Slide 6: Operation Modes
Conductive Atomic Force Microscopy | Wiley Online Books

## Slide 7: Choice of Probe
Tips using a graphene or diamond coating are much more robust and sensitive
Highly robust
Demands tedious fabrication, but can be made in bulk
Conductive varnish on insulating probe (often, Si)
Optimize varnish thickness against tip diameter
Tips prone to extremely short lifespan
Highly etched & lithographically structured probes (thin films or wire)
Allows fine control over structure and conductivity
Exceedingly difficult to manufacture, sometimes

## Slide 8: Reliability of Conductive AFM Probes
Conductive Atomic Force Microscopy | Wiley Online Books
M. Lanza, A. Bayerl, T. Gao, M. Porti, M. Nafria, G. Y. Jing, Y. F. Zhang, Z. F. Liu, and H. L. Duan, Advanced Materials 25, 1440 (2013).
Metal-varnished tip
Etched & Wire tips
Diamond & Graphene Coated Tips
Pt-Ir Tip
Used Pt-Ir Tip

## Slide 9: Applications of C-AFM
Year | Achievements
2011 | Monolayer MoS₂ n-type FET with a high-k top gate dielectric
 | Integrated logic circuit based on a single-layer MoS₂
2012 | Monolayer WSe₂ p-type FET with a high-k top gate dielectric
 | Chemical vapor deposition (CVD) of MoS₂ atomic layers on SiO₂ substrates
2014 | FET built from all 2D material components
 | Vertical and in-plane heterostructures of WS₂/MoS₂ monolayers grown by CVD
 | Atomic layer deposition (ALD) of MoS₂ thin films
2015 | CVD of epitaxial monolayer MoS₂ on sapphire
 | Metal organic CVD of a highly uniform monolayer MoS₂ on 100 mm SiO₂ wafers
 | WSe₂-based CMOS and integrated circuits
 | Tunnel diodes based on MoS₂/WSe₂ vdW heterostructures grown by CVD
2016 | MoS₂ FET with a 1 nm gate length
 | MoS₂ FET with a sub −10 nm channel length
 | CVD of epitaxial monolayer MoS₂ on gallium nitride
 | Pulsed laser deposition (PLD) of MoS₂ thin films
2017 | Ambipolar MoS₂ FET by tailoring the Schottky barrier with oxygen plasma
 | Ohmic contacts to monolayer MoS₂ by van der Waals bonded metal/h-BN electrodes
 | Demonstration of a microprocessor based on CVD grown bilayer MoS₂
 | Hot electron transistor with GaN emitter, graphene base, and WSe₂ base-collector barrier
2018 | Batch CVD growth of uniform monolayer MoS₂ on 150 mm soda-lime glass wafers
2019 | Van der Waals contacts between 3D metals and 2D semiconducting TMDs
(ACS 2017) Shottky Barrier Tuned via Plasma Treatment in 2D MoS₂
Schottky barrier height maps and histograms on pristine MoS2 (a,d) and after O2 plasma for 300 (b,e) and 600 s

## Slide 10: Final Notes on C-AFM: Atomic Resolution Possible (sometimes)
Note- current is dependent on the effective area between the tip and sample
Insulating samples are therefore easier to probe at high resolution
S. A. Sumaiya and M. Z. Baykara, True Atomic-Resolution Imaging under Ambient Conditions via Conductive Atomic Force Microscopy, (n.d.).

## Slide 11: Development of Carbon Nanotubes
(1991)
(1993)
(today)
Multi-walled carbon nanotubes (MWCNT) become a “hot” research area
First single-walled carbon nanotubes (SWCNT) fabricated
(2000-2001)
Experimental characterization of structural, thermal, electronic properties
Commercial use: composites, high lubricity, thermal interfaces
CNT Sensors
Quantum wire applications
Li-ion batteries

## Slide 12: Fact Sheet
Can be made highly conductive or semiconducting
  SWCNT conductance up to ~10⁷ A/mm²
  Copper conductance around ~5.8 x 10⁷ A/mm²
Can be engineered by chirality, defects, doping
Highly Tunable Electronic Modes
Y. Xiang, L. Zhang, and C. Liu, Nanomaterials 15, 1165 (2025).
Synthesis has been developed for 30+ years; not challenging to grow or purchase
Enables probes with tip diameter <15 nm
Extreme Aspect Ratio (>0.1ᴇ⁹:1)
Y. Xiang, L. Zhang, and C. Liu, Nanomaterials 15, 1165 (2025).
Cylinders buckle when stressed
Conductivity not lost when this happens
Structural Integrity
H. Shima, Buckling of Carbon Nanotubes, Materials 5, 47 (2012).

## Slide 13: Carbon Nanotubes as C-AFM Probes
Attach via micromanipulator and some carbon paste
Cut with FIB for greater integrity
Weld to probe using electron beam deposition
*there exist other methods for attachment, which can require less equipment
A. D. Slattery, C. J. Shearer, C. T. Gibson, J. G. Shapter, D. A. Lewis, and A. J. Stapleton, Nanotechnology 27, 475708 (2016).

## Slide 14: Carbon nanotubes as a path to overcome traditional C-AFM limitations
Stable, sensitive conductance
Comparison in terms of imaging performance on a sample of highly ordered pyrolytic graphite (HOPG).
Improved contrast over typical Pt-Ir tips
  CNTs more conductive than most materials
Signal integrity unlikely to degrade compared to other tips
A. D. Slattery, C. J. Shearer, C. T. Gibson, J. G. Shapter, D. A. Lewis, and A. J. Stapleton, Nanotechnology 27, 475708 (2016).

## Slide 15: Carbon nanotubes as a path to overcome traditional C-AFM limitations
2. Structural Integrity and Repairability
The authors use a small bundle of SWCNTs
  Slightly larger diameter but much better resilience
Authors note CNTs are prone to buckling instead of chipping
CNT-tips offer the ability to repair a tip

## Slide 16: Carbon nanotubes as a path to overcome traditional C-AFM limitations
3. Improved tip diameter
May or may not improve spatial resolution of C-AFM very much (depends on sample)
This does give big improvements to topology map (which can be gathered simultaneously)
  Necessary for characterization of nanoelectronics
J. Xu, Y. Shingaya, Y. Zhao, and T. Nakayama, Applied Surface Science 335, 11 (2015).
Au-coated Si Tip
CNT Tip

## Slide 17: Potential research directions
Little work has been done using such tips explicitly yet (I think)
Such long-duration experiments (such as extended chemical reactions) cannot be imaged using most C-AFM probes
  (Hundreds of scans required)
V. Jain*, S. Ahsanullah*, G. Saggars*, A. Mathai*, D. Pates, A. Guardiola, K. Watanabe, T. Taniguchi, M. A. McGuire, D. Graf, M. Gibertini, A. F. May, D. Ovchinnikov, “Electrical detection of magnetic states in G-type van der Waals antiferromagnet”, In Preparation (2025)

## Slide 18: That’s all! grant-saggars@ku.edu November 4, 2025